{"text_format":"2","workbook":"xlsx","n_sheets":2,"sheets":["Final","Sheet1"],"not_rendered":{"charts":0,"images":0,"macros":0,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"run_cost":"one","rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Final","index":0,"n_rows":28,"n_cols":6,"rows_kept":28,"cut":false,"hidden":false,"merged":[[26,4,29,4],[26,5,29,5],[2,5,7,5],[2,1,7,1],[18,1,24,1],[8,1,17,1],[18,5,24,5],[8,5,17,5]]}
[0,null,"Are all pressure and flow axes present and is each labelled and with timescale, and are the filled and voided volumes data presented in some form ? ",null,null,"Screening"]
[null,null,"Question","Y / N / na","Reviewer comment"]
[1,"Pre-filling","Were all of the pressure axes displayed with the same height per cmH2O?",null,null,"Pre-filling"]
[2,null,"Was the urine flow vertical axis size as recommended (i.e. 0-25 up to 0-50 ml/sec)?"]
[3,null,"Were the pves and the pabd marked on the trace as being zeroed to atmosphere?"]
[4,null,"Was a good quality* cough test carried out at the very start of the test?"]
[5,null,"Were the initial resting pves and pabd pressures in the physiological range**? "]
[6,null,"Was initial resting pdet in the physiological range**?"]
[7,"Filling","Were cough tests or Valsalvas visible on the printout during filling?",null,null,"Filling"]
[8,null,"Was each set of cough peaks or Valsalvas good quality* throughout (or after correction)?"]
[9,null,"Were small 'live' pressure fluctuations (e.g. patient breathing) visible throughout the test (or after any correction) equally on pabd and pves, but not visible on pdet?"]
[10,null,"If there was tube leakage (steady pressure descent) was it corrected?"]
[11,null,"Was the patient position recorded on the trace at any point?"]
[12,null,"If patient position change was evident on the trace, was the transducer level adjusted?"]
[13,null,"If detrusor overactivity was present, was it clearly marked at any point?"]
[14,null,"If poor compliance was seen (pves rising continuously), was the pump stopped for some time?"]
[15,null,"Were the patient reported sensations of filling indicated at any point?"]
[16,null,"If the flow trace shows urine leakage was present, was it clearly marked as such?"]
[17,"Voiding","Was a good quality* cough test done at the end of filling, before voiding?",null,null,"Voiding"]
[18,null,"Do all traces remain in view during very high or low pressures?"]
[19,null,"Was either ‘permission to void’ or 'void' indicated?"]
[20,null,"Were the markers for start and end of void correctly placed?"]
[21,null,"Was a good quality* cough test done after the final void?"]
[22,null,"Was the Qmax marker placed on the flow trace?"]
[23,null,"Was the Qmax marker placed away from artefacts?"]
[null,null,"Quality Score:",{"$f":"(COUNTIF(D3:D25,\"Y\")/(COUNTIF(D3:D25,\"Y\")+COUNTIF(D3:D25,\"N\")+COUNTIF(D3:D25,\"\")))","v":0}]
["* Note:",null,"Good quality cough is one where the smaller spike is greater than 70% of the larger spike",null,null,"Further comments"]
["** Note: ",null,"Physiological range for initial resting pressures\npdet: -5 to +5 cmH2O\npves and pabd: 5–20 cmH2O if supine, 15–40 cmH2O if seated, 30–50 cmH2O if standing"]

{"sheet":"Sheet1","index":1,"n_rows":3,"n_cols":1,"rows_kept":3,"cut":false,"hidden":false,"merged":[]}
["Y"]
["N"]
["na"]